ADMARNEURO

LAPTOP SETUP MANUAL

Clone Drive Creation & System Restoration Procedures

Supported Systems Cadwell Easy 2/3 • Sierra Summit • Sierra Wave Sierra Wedge • Plinth • Natus PIU/VikingQuest/UltraPro

Technical Support: (561) 964-0477

Email: tech@admarneuro.com

TABLE OF CONTENTS

| Section 1: Quick Reference | ......................... 1 |
| --- | --- |
| 1.1 Setup Methods Overview |  |
| 1.2 Hardware Requirements |  |
| Section 2: Windows 7 Backup Method | ......................... 2 |
| 2.1 Prerequisites |  |
| 2.2 Creating a Clone Drive |  |
| 2.3 Restoring from Clone Drive |  |
| Section 3: Rescuezilla Method (Preferred) | ......................... 3 |
| 3.1 Prerequisites |  |
| 3.2 Creating a Clone Drive |  |
| 3.3 Restoring from Clone Drive |  |
| Section 4: Sierra Wedge Setup | ......................... 4 |
| 4.1 Windows Installation |  |
| 4.2 Sierra Wedge Software Installation |  |
| 4.3 Network Configuration |  |
| 4.4 Connection Verification |  |
| Section 5: Sierra Wave Setup | ......................... 5 |
| 5.1 Windows 11 Pro Preparation |  |
| 5.2 Office Installation |  |
| 5.3 Prerequisites Installation |  |
| 5.4 Sierra Wave 11.0 Installation |  |
| Section 6: Natus Systems Setup | ......................... 7 |
| 6.1 System Requirements |  |
| 6.2 Windows 11 Pro Upgrade |  |
| 6.3 Office 2016 Installation |  |
| 6.4 Natus Synergy & Viking Installation |  |
| 6.5 Remote Utilities Installation |  |
| 6.6 PC Promise Confirmation |  |
| Section 7: Network Configuration | ......................... 9 |
| 7.1 Adapter Settings |  |
| 7.2 IP Address Configuration |  |
| Section 8: Glossary of Terms | ......................... 10 |

# Section 1: Quick Reference

## 1.1 Setup Methods Overview

Use the method corresponding to your specific hardware system.

| METHOD 1 | METHOD 2 ✓ | METHOD 3 | METHOD 4 |
| --- | --- | --- | --- |
| Windows 7 Backup | Rescuezilla (Preferred) | USB | USB & PC Promise |
| Scope: Single System Systems: • Cadwell Easy 2 • Cadwell Easy 3 • Sierra Summit See Section 2 | Scope: Single & Combined Systems: • Plinth • Easy 2 & Wave See Section 3 | Scope: Single System Systems: • Sierra Wave • Sierra Wedge See Sections 4-5 | Scope: Natus Systems Systems: • Natus PIU • Natus VikingQuest • Natus UltraPro See Section 6 |

## 1.2 Hardware Requirements

| SYSTEM TYPE | RAM | PROCESSOR | HARD DRIVE | OS VERSION |
| --- | --- | --- | --- | --- |
| Natus (PIU, VikingQuest, UltraPro, EDX) | 8 GB | Intel i3/i5/i7 | 256 GB | Windows 11 Pro |
| Sierra Wave | 8 GB | Intel i3/i5/i7 | 256 GB | Windows 11 Pro |
| Combo System (Wave + Other) | 8 GB | — | 500 GB+ | Windows 11 Pro |
| Sierra Wedge | 4 GB | — | 256 GB | Win 10 (32-bit) |
| Plinth | 4 GB | — | 256 GB | Windows 10 Pro |

# Section 2: Windows 7 Backup Method

For: Cadwell Easy 2, Cadwell Easy 3, Sierra Summit

## 2.1 Prerequisites

•  External USB hard drive with sufficient storage

•  Similar laptop (must be in the same series)

•  Windows Installation/Repair USB drive

## 2.2 Creating a Clone Drive

| Step | Action |
| --- | --- |
| 1 | Connect external USB hard drive to source laptop |
| 2 | Navigate to: Control Panel → System and Security → Backup and Restore (Windows 7) |
| 3 | Click "Create a system image" on left sidebar |
| 4 | Select "On a hard disk" and choose your USB drive from dropdown → Click Next |
| 5 | Confirm partitions are checked (C: and System Reserved/EFI) → Click Next |
| 6 | Click Start backup — wait for progress bar to complete |
| 7 | Skip System Repair Disc prompt (if you have installation media). Safely eject drive. |

## 2.3 Restoring from Clone Drive

| Step | Action |
| --- | --- |
| 1 | Insert Windows Installation/Repair USB into target laptop |
| 2 | Boot from USB (press F12 or Del at startup to access boot menu) |
| 3 | On Windows Setup screen: Click Next → Select "Repair your computer" (NOT "Install Now") |
| 4 | Navigate to: Troubleshoot → Advanced Options → System Image Recovery |
| 5 | Connect external drive with backup — system auto-detects image → Select and click Next |
| 6 | ⚠ Check "Format and repartition disks" → Click Finish → Yes (THIS ERASES ALL DATA) |
| 7 | Wait for restore to complete — laptop restarts into cloned environment |

# Section 3: Rescuezilla Method (Preferred)

For: Plinth, Easy 2 & Wave Combo Systems

## 3.1 Prerequisites

•  Rescuezilla bootable USB drive

•  External storage drive for backup image

•  Laptop meeting minimum system requirements

## 3.2 Creating a Clone Drive

| Step | Action |
| --- | --- |
| 1 | Insert Rescuezilla USB into source laptop |
| 2 | Restart and enter Boot Menu (F12, F9, or Del depending on manufacturer) |
| 3 | Select USB drive to boot Rescuezilla |
| 4 | Select language — wait for graphical interface to load |
| 5 | Click the Backup icon |
| 6 | Select laptop's internal hard drive as source → Click Next |
| 7 | Connect external storage drive — select as destination → Click Next |
| 8 | Name the backup (e.g., "Sierra_Summit_Master_2026-02-02") |
| 9 | Review summary → Click Next to begin backup |
| 10 | Wait for "Backup Completed Successfully" message |
| 11 | Shut down laptop and remove USB drives |

## 3.3 Restoring from Clone Drive

| Step | Action |
| --- | --- |
| 1 | Insert Rescuezilla USB + external drive (with image) into target laptop |
| 2 | Boot from Rescuezilla USB |
| 3 | Select language to load interface |
| 4 | Click the Restore icon |
| 5 | Choose external drive — navigate to image folder and select it → Click Next |
| 6 | ⚠ Select target laptop's internal drive as destination (THIS ERASES ALL DATA) |
| 7 | Confirm drive details → Click Next → Click Yes to start restore |
| 8 | Wait for "Restore Completed Successfully" message |
| 9 | Shut down, remove USB drives, and power on to verify system boots correctly |

# Section 4: Sierra Wedge Setup

USB Method — Requires Windows 10 32-bit

## 4.1 Windows Installation

⚠ CRITICAL: Sierra Wedge requires Windows 10 32-bit ONLY

### Prepare Installation Media

1.  Use Microsoft Media Creation Utility

2.  Create bootable flash drive with Windows 10 32-bit

### Install Windows

| Step | Action |
| --- | --- |
| 1 | Insert Windows 10 32-bit USB drive |
| 2 | Spam F12 + FN keys until boot screen appears |
| 3 | Select USB Storage from boot menu |
| 4 | ⚠ Begin Windows setup — DO NOT CONNECT TO INTERNET |
| 5 | Delete all existing partitions during setup |
| 6 | Complete Windows setup |

## 4.2 Sierra Wedge Software Installation

| Step | Action |
| --- | --- |
| 1 | Insert Wedge USB drive |
| 2 | Open File Explorer → Navigate to USB drive (usually F:) |
| 3 | Open "Wedge Disk" folder |
| 4 | Run Setup (file with Sierra icon) |
| 5 | Complete Sierra XP setup |

## 4.3 Network Configuration

Required for base unit communication. See Section 7 for detailed steps.

## 4.4 Connection Verification

1.  Connect laptop to base unit

2.  Navigate to: File Explorer → This PC → Local Disk (C:) → Cadwell folder

3.  Search for "Sierra IP"

4.  Confirm "Registry Found" message displays

✓ If "Registry Found" appears, base connection is verified

# Section 5: Sierra Wave Setup

USB Method — Requires Windows 11 Pro 64-bit

## 5.1 Windows 11 Pro Preparation

### Internet Bypass (Skip if Windows 11 Pro is preloaded)

| Step | Action |
| --- | --- |
| 1 | During Windows setup, press Shift + F10 to open Command Prompt |
| 2 | Type: oobe\bypassnro → Press Enter |
| 3 | System will automatically reboot |
| 4 | Select "I do not have internet" when prompted |

⚠ If bypass fails, reinstall Windows from scratch

### Disable Automatic Updates (Critical)

Windows updates can break medical device software. Disable them immediately.

| FOR WINDOWS 11 PRO | FOR NON-PRO WINDOWS |
| --- | --- |
| 1. Press Win + R, type gpedit.msc, Enter 2. Navigate to: Computer Configuration → Administrative Templates → Windows Components → Windows Update → Manage updates offered 3. Double-click "Do not include drivers with Windows Update" 4. Select Enabled → Apply → OK 5. Go to: ...Windows Update → Manage end user experience 6. Double-click "Configure Automatic Updates" 7. Select Disabled → Apply → OK | Group Policy Editor is not available on non-Pro editions. Alternative: 1. Download Windows Update Blocker v1.8 2. Source: sordum.org 3. Run the utility 4. Select "Disable Updates" 5. Click "Apply Now" |

## 5.2 Office Installation

### Remove Office Trial (Required First)

1.  Save office removal script as office_removal.cmd

2.  Right-click → Run as Administrator

3.  Wait for "Office Removal Complete" message

4.  Restart computer

### Install Office 2007

1.  Insert Office thumb drive

2.  Run “OfficeSetup 64 bit”

3.  Enter serial number when prompted - WRJBM-2MYC9-V4BG9-H7JPM-2BVD8

4.  Complete installation

## 5.3 Prerequisites Installation

| Step | Action |
| --- | --- |
| 1 | Configure network adapter (see Section 7 for detailed steps) |
| 2 | Reboot and connect to internet |
| 3 | Go to: Control Panel → Programs → Programs and Features → Turn Windows Features on or off |
| 4 | Check .NET Framework 3.5 → Click OK → Install via Windows Update |
| 5 | Reboot |
| 6 | Go to: Control Panel → System and Security → Security and Maintenance → Change User Account Control Settings |
| 7 | From (wave) thumb drive "Special" folder: Install Microsoft Visual C++ Redistributables (vc_redist.x64) |
| 8 | Create folder named "Admar" on C: drive |
| 9 | Install DirectX 9+ (if prompted to browse, select "Admar" folder) |

## 5.4 Sierra Wave 11.0 Installation

| Step | Action |
| --- | --- |
| 1 | Click Install QuickReport → Follow prompts |
| 2 | Restart when prompted |
| 3 | Run autorun from Sierra Wave 11.0 installerClick Install QuickReport → Follow prompts |
| 4 | Select "Yes" to remove old QuickReport files (templates are preserved) |
| 5 | Select "Yes to All" for shared files prompt |
| 6 | Click OK when uninstall completes → Click Next (default: C:\Cadwell\QuickReport) |
| 7 | Click Install Sierra Wave |
| 8 | Click Update Help Files |
| 9 | Click Update Splash Screen |
| 10 | Close installer, eject media, restart computer |

# Section 6: Natus Systems Setup

USB & PC Promise Method — For PIU, VikingQuest, UltraPro

## 6.1 System Requirements

| Operating System | Windows 11 Pro (64-bit) |
| --- | --- |
| RAM | 8 GB minimum |
| Processor | Intel i Series (i3, i5, or i7) |
| Hard Drive | 256 GB minimum |

## 6.2 Windows 11 Pro Upgrade (If Needed)

1.  Check current edition: Settings → System → About

2.  If not Windows 11 Pro, go to: Settings → System → Activation

3.  Select "Upgrade your edition of Windows"

4.  Enter Windows 11 Pro product key

5.  Complete upgrade and restart

## 6.3 Office 2016 Installation

1.  Remove any Office trial versions (run office_removal.cmd as Administrator)

2.  Purchase Office 2016 License Key from Product Keys vendor

3.  Insert "OFFICEWS" thumb drive

4.  Run installer from thumb drive

5.  Enter product key when prompted

6.  Complete activation

## 6.4 Natus Synergy & Viking Installation

1.  Download Natus Synergy & Viking software

2.  Run installer as Administrator

3.  Follow installation prompts

4.  When prompted for License ID, enter using the format below:

NATUS LICENSE ID FORMAT TA + [Year] + [4 digits] + M Example: TA260004M (Year 26, Unit #0004)

5.  Complete installation and restart

## 6.5 Remote Utilities Installation

| Step | Action |
| --- | --- |
| 1 | Go to: remoteutilities.com/download |
| 2 | ⚠ Download HOST (not Viewer) |
| 3 | Right-click installer → Run as Administrator |
| 4 | Accept defaults and complete installation |
| 5 | Note your Internet-ID (9 digits displayed as XXX-XXX-XXX) |
| 6 | Click Settings (gear icon) → Authentication tab → Enable Single Password → Enter password → OK |

## 6.6 PC Promise Confirmation (Required)

⚠ You must email PC Promise to confirm license activation

Include in your email to PC Promise: • License ID (e.g., TA260004M) • Software (Synergy or Viking) • System Type (PIU, VikingQuest, or UltraPro) • Remote Utilities Internet-ID (XXX-XXX-XXX) • Password (synergy!)

# Section 7: Network Configuration

Required for Sierra Wedge and Sierra Wave base unit communication

## 7.1 Adapter Settings

| Step | Action |
| --- | --- |
| 1 | Go to: Control Panel → Network and Sharing Center |
| 2 | Click "Change adapter settings" (left sidebar) |
| 3 | Right-click Ethernet → Properties |
| 4 | Click Configure button |
| 5 | Go to Advanced tab |
| 6 | Scroll to "Speed and Duplex" → Value → Select "10 Mbps Full Duplex" → OK |

## 7.2 IP Address Configuration

| Step | Action |
| --- | --- |
| 1 | In Ethernet Properties, double-click "Internet Protocol Version 4 (TCP/IPv4)" |
| 2 | Click Advanced button |
| 3 | Go to WINS tab → Select "Disable NETBIOS over TCP/IP" |
| 4 | Click OK to close Advanced settings |
| 5 | Select "Use the following IP address" |
| 6 | Enter the following: IP Address: 192.168.113.66 Subnet Mask: 255.255.255.0 (auto-fills) |
| 7 | Click OK |
| 8 | ⚠ Uncheck ALL items EXCEPT "Internet Protocol Version 4 (TCP/IPv4)" |
| 9 | Click OK to close Ethernet Properties |

# Section 8: Glossary of Terms

| Term | Definition |
| --- | --- |
| Base Unit | The main hardware component of Sierra systems that connects to the laptop via Ethernet and performs signal acquisition. |
| Boot Menu | BIOS/UEFI menu accessed during startup (F12, F9, or Del) that allows selection of boot device. |
| Clone Drive | An exact copy of a system's hard drive, including OS, software, and configurations, used for rapid deployment. |
| CMOS Battery | Small battery on motherboard that maintains BIOS settings when computer is powered off. |
| DirectX | Microsoft multimedia API required by some medical software for graphics and input device handling. |
| EEG | Electroencephalography — measures electrical activity in the brain using scalp electrodes. |
| EMG | Electromyography — measures electrical activity produced by skeletal muscles. |
| gpedit.msc | Group Policy Editor — Windows Pro tool for configuring system policies including update settings. |
| Internet-ID | Unique 9-digit identifier (XXX-XXX-XXX) assigned by Remote Utilities for remote access connections. |
| IP Address | Numerical label assigned to devices on a network. Sierra systems use static IP 192.168.113.66. |
| .NET Framework | Microsoft software framework required by many Windows applications, including Sierra Wave. |
| NETBIOS | Legacy network protocol — disabled for Sierra systems to prevent communication conflicts. |
| OOBE | Out-of-Box Experience — Windows initial setup screen that runs on first boot. |
| Partition | Logical division of a hard drive. System images typically include C: and System Reserved partitions. |
| PC Promise | Third-party licensing service for Natus systems — must be contacted to confirm license activation. |
| PIU | Patient Interface Unit — Natus hardware component for patient signal acquisition. |
| QuickReport | Cadwell report generation software included with Sierra Wave installations. |
| Remote Utilities | Third-party remote access software used by AdmarNeuro for technical support sessions. |
| Rescuezilla | Open-source disk imaging tool that operates outside Windows for creating and restoring system clones. |
| Sierra Summit | Cadwell neurodiagnostic system for EMG, NCS, and EP studies. |
| Sierra Wave | Cadwell advanced neurodiagnostic system with expanded capabilities. |
| Sierra Wedge | Cadwell portable neurodiagnostic system requiring Windows 10 32-bit. |
| Subnet Mask | Network parameter defining IP address range. Sierra systems use 255.255.255.0. |
| System Image | Complete backup of a computer's hard drive including OS, applications, and settings. |
| TCD | Transcranial Doppler — ultrasound technique measuring blood flow velocity in brain arteries. |
| UAC | User Account Control — Windows security feature controlling administrative privileges. |
| vc_redist | Visual C++ Redistributable — Microsoft runtime libraries required by many applications. |
| VikingQuest | Natus EMG/NCS neurodiagnostic system. |
| WINS | Windows Internet Name Service — legacy name resolution disabled for Sierra network configuration. |